EKG Embedded System Troubleshooting Log

Purpose

This document tracks real problems encountered during development of the Raspberry Pi–based EKG system, along with root causes and implemented solutions. It serves as both a debugging reference and a foundation for future design improvements and research documentation.

1. AD8232 Power Verification Failure

Problem
The AD8232 showed no LED activity, making it unclear whether the chip was powered.

Observed Behavior

- No blinking LED

- No measurable output signal

Root Cause
Power delivery through jumper wires was assumed but not verified. Breadboard connections can be unreliable, especially with loose Dupont connectors.

Solution

- Used a multimeter to probe VCC and GND directly at the chip pins

- Confirmed ~3.3 V supply

- Ensured all grounds (Pi, ADC, AD8232) shared a common reference

Takeaway
Always verify power electrically, not visually. LEDs are not sufficient indicators of proper operation.

2. Incorrect or Non-Physiological ECG Waveform

Problem
The output waveform did not resemble a recognizable ECG signal.

Observed Behavior

- Irregular spikes

- No clear P-QRS-T structure

Root Cause

- Improper electrode placement

- Incorrect lead configuration

- Poor skin contact or inconsistent grounding

Solution

- Standardized electrode placement (RA(blue), LA(red), RL(black) configuration)

- Ensured firm adhesion using clinical-grade electrodes

- Verified continuity through the 3.5 mm lead cable

Takeaway
Signal acquisition quality is dominated by electrode placement and contact.

3. Constant Lead-Off Detection

Problem
The system continuously reported lead-off conditions even when electrodes were attached.

Observed Behavior

- LO+ / LO− signals triggered constantly

- No stable ECG output

Root Cause

- GPIO pins not properly configured

- Weak or inconsistent electrode contact

Solution

- Configured GPIO pins with internal pull-ups

- Verified wiring from AD8232 LO pins to Raspberry Pi GPIO

- Replaced or repositioned electrodes to improve contact

Takeaway
Lead-off detection is sensitive to both electrical configuration and physical contact quality.

4. Excessive Noise and Motion Artifact

Problem
The ECG signal was highly noisy and degraded significantly with movement.

Observed Behavior

- High-frequency jitter

- Large baseline shifts during motion

- Signal distortion during patient movement

Root Cause

- Motion-induced electrode impedance changes

- Environmental electrical noise

- Lack of filtering pipeline

Solution

- Implemented software filtering:

  - Baseline removal (high-pass behavior)

  - Smoothing filter (low-pass behavior)

- Improved wire stability and strain relief

- Identified IMU-based correction as a future enhancement

Takeaway
Motion artifact is a dominant challenge in real-world ECG systems and must be addressed at both hardware and software levels.

5. Flatline Signal Due to SDN Pin Misconfiguration

Problem
The ECG signal appeared flat (no waveform) but would produce spikes when the leads were shorted or disturbed.

Observed Behavior

- Constant flatline output

- Spikes only when leads were touched or shorted

Root Cause
The SDN (shutdown) pin on the AD8232 was not tied to 3.3 V. This kept the chip in a disabled or inactive state, preventing normal signal amplification.

Solution

- Tied the SDN pin directly to 3.3 V

- Verified the chip exited shutdown mode

- Confirmed normal ECG waveform appeared after correction

Takeaway
Critical control pins like SDN must be explicitly set. A floating or low SDN pin will disable the analog front end and produce misleading flatline behavior.

6. ADS1115 Not Producing Expected Data

Problem
ADC readings were inconsistent or not appearing as expected.

Observed Behavior

- No data or incorrect values

Root Cause

- I2C misconfiguration

- Incorrect channel selection

Solution

- Verified device presence using i2cdetect -y 1

- Confirmed correct I2C address (typically 0x48)

- Used correct channel index (0 for A0)

Takeaway
Always verify communication layer before debugging signal quality.

7. ADS1115 Library Attribute Error

Problem
Error: ADS1115 has no attribute P0

Observed Behavior

- Program crashes at channel initialization

Root Cause
Library API differences or incorrect usage

Solution

- Replaced symbolic constant with integer channel index:

  - AnalogIn(adc_device, 0)

Takeaway
Library abstractions can vary. Always confirm expected API usage.

8. Raw ADC Output Appeared Random

Problem
Terminal output appeared as random or meaningless values.

Observed Behavior

- Rapidly changing numbers

- No obvious waveform structure

Root Cause

- Raw ADC values were being printed without scaling or filtering

- Noise present in signal

Solution

- Interpreted values as sampled voltage

- Added filtering and visualization (plotting)

Takeaway
Raw biosignals are not human-readable without processing.

9. Python Environment and Dependency Issues

Problem
Errors occurred when installing required libraries such as busio.

Observed Behavior

- pip install failures

- "externally managed environment" errors

Root Cause
System Python restrictions on Raspberry Pi OS

Solution

- Created a Python virtual environment

- Installed all dependencies within the virtual environment

Takeaway
Use isolated environments for embedded Python development to avoid system conflicts.

10. Baseline Drift in ECG Signal

Problem
The ECG signal slowly drifted upward or downward over time.

Observed Behavior

- Signal baseline not centered

- Long-term trend shifts

Root Cause

- Electrode polarization

- Low-frequency physiological and environmental noise

Solution

- Implemented baseline tracking using a slow-moving average

- Subtracted this baseline from the raw signal

Takeaway
Baseline drift must be removed for meaningful ECG interpretation.

11. High-Frequency Noise and Jitter

Problem
The signal contained significant high-frequency fluctuations.

Observed Behavior

- Jagged waveform

- Poor readability of cardiac features

Root Cause

- Electrical interference

- ADC quantization noise

Solution

- Applied smoothing filter (low-pass behavior)

- Tuned smoothing parameter for balance between clarity and responsiveness

Takeaway
Filtering is required to extract clinically useful information from raw signals.

12. Lack of Real-Time Filtering Pipeline

Problem
Initial system displayed only raw ECG data.

Observed Behavior

- Noisy, unstable waveform

Root Cause
No signal processing pipeline implemented

Solution

- Built real-time filtering chain:

  - Baseline removal (high-pass behavior)

  - Smoothing filter (low-pass behavior)

Takeaway
A minimal signal processing pipeline is required for any usable ECG display.

Future Additions

- IMU-based motion artifact correction

- Hardware filtering (RC filters, shielding)

- Improved electrode interface design

- Real-time feature extraction (QRS detection)

End of document